## Instructions
Objectives | Instructions
The Joint Commission is internationally recognized by the medical industry and government agencies for performance standards governance designed to enhance a hospital's operational effectiveness and efficiency. Conducting a gap analysis is a key step in developing and implementing a dose management program to comply with the latest Joint Commission requirements. The objectives of this analysis include: · Assisting select department and facilities in determining how current management programs and practices align with recent Joint Commission Prepublication Standards · Devising a plan of corrective action based on identifiable performance gaps to increase compliance with applicable Joint Commission requirements and other relevant standards. | 1. Identify and assign an owner to each requirement. 2. Have each owner evaluate the existing practices, summarize findings/supporting information and assess current status of the requirement. 3. Based on the assessment, determine whether your organization has met, partially met, or not met the revised requirement, or whether this requirement is not applicable (N/A) for your institution. Indicate the result by filling in the ‘Status’ column on the Gap Analysis Tool with an "x". 4. Identify and describe an associated Corrective Action for each requirement that is partially met or not met. 5. Implement Corrective Actions to close the ‘gap’ in your operational performance. 6. Generate a sample executive summary report under the Summary tab.
Gap Analysis
This gap analysis is a self-evaluation. The list of Joint Commission requirements are presented in a manner that will help you compare your organization’s current practices against the recent Joint Commission Prepublication Requirements for Diagnostic Imaging Services (issued January 9, 2015, effective July 1, 2015). This comparison results in a baseline assessment of your organization and helps you to identify current programs that are in place for managing issues, including existing policies and procedures, roles and responsibilities, communications, training programs, and others.
As you conduct the analysis, you will recognize strengths and weaknesses of your organization’s present operations and determine the extent to which your operations conform to the Joint Commission Prepublication Requirements.
Upon completion of the analysis, you should be able to identify which processes are already in place and working well, and which you need to develop or further refine. Conducting the gap analysis will also help you become familiarized with the Joint Commission Prepublication Requirements and the elements critical to developing, implementing, checking, and promoting adherence on an on-going basis.

## Gap Analysis
The Joint Commission Revised Standards Gap Analysis
| Requirements | Remark | Affected Modality | Owner | Finding / Supporting Information | Status (fill in "x") |  |  |  | Corrective Action
[6 empty] | Not in place | Partially met | Fully met | N/A
Environment of Care (EC) | [5 empty] | x
EC.02.01.01 | The [critical access] hospital manages safety and security risks.
A 14 | The [critical access] hospital manages magnetic resonance imaging (MRI) safety risks associated with the following: • Patients who may experience claustrophobia, anxiety, or emotional distress • Patients who may require urgent or emergent medical care • Patients with medical implants, devices, or imbedded metallic foreign objects (such as shrapnel) • Ferromagnetic objects entering the MRI environment • Acoustic noise |  | MRI
A 16 | The [critical access] hospital manages magnetic resonance imaging (MRI) safety risks by doing the following: • Restricting access of everyone not trained in MRI safety or screened by staff trained in MRI safety from the scanner room and the area that immediately precedes the entrance to the MRI scanner room. • Making sure that these restricted areas are controlled by and under the direct supervision of staff trained in MRI safety. • Posting signage at the entrance to the MRI scanner room that conveys that potentially dangerous magnetic fields are present in the room. Signage should also indicate that the magnet is always on except in cases where the MRI system, by its design, can have its magnetic field routinely turned on and off by the operator. | 3: Direct impact requirements apply | MRI
EC.02.02.01 | The [critical access] hospital manages risks related to hazardous materials and waste.
A 17 | For [critical access] hospitals that provide computed tomography (CT), positron emission tomography (PET), or nuclear medicine (NM) services: The results of staff dosimetry monitoring are reviewed at least quarterly by the radiation safety officer, diagnostic medical physicist, or health physicist to assess whether staff radiation exposure levels are “As Low As Reasonably Achievable” (ALARA) and below regulatory limits. 1 2 |  | CT, PET, NM
EC.02.04.01 | The [critical access] hospital manages medical equipment risks.
A 10 | The [critical access] hospital identifies quality control and maintenance activities to maintain the quality of the diagnostic computed tomography (CT), positron emission tomography (PET), magnetic resonance imaging (MRI), and nuclear medicine (NM) images produced. The [critical access] hospital identifies how often these activities should be conducted. (See also EC.02.04.03, EP 15). |  | CT, PET, MRI, NM
EC.02.04.03 | The [critical access] hospital inspects, tests, and maintains medical equipment.
C 15 | The [critical access] hospital maintains the quality of the diagnostic computed tomography (CT), positron emission tomography (PET), magnetic resonance imaging (MRI), and nuclear medicine (NM) images produced. (See also EC.02.04.01, EP 7) |  | CT, PET, MRI, NM
A 17 | For diagnostic computed tomography (CT) services: At least annually, a diagnostic medical physicist does the following: 2 3 4 • Measures the radiation dose (in the form of volume computed tomography dose index [CTDIvol]) produced by each diagnostic CT imaging system for the following four CT protocols: o adult brain, o adult abdomen, o pediatric brain, and o pediatric abdomen. If one or more of these protocols is not used by the [critical access] hospital, other commonly used CT protocols may be substituted. • Verifies that the radiation dose (in the form of CTDIvol) produced and measured for each protocol tested is within 20 percent of the CTDIvol displayed on the CT console. The dates, results, and verifications of these measurements are documented. | D: Documentation is required; 3: Direct impact requirements apply | CT
A 19 | For diagnostic computed tomography (CT) services: At least annually, a diagnostic medical physicist conducts a performance evaluation of all CT imaging equipment. The evaluation results, along with recommendations for correcting any problems identified, are documented. The evaluation includes the use of phantoms to assess the following imaging metrics: 2 4 • Image uniformity • Slice thickness accuracy • Slice position accuracy (when prescribed from a scout image) • Alignment light accuracy • Table travel accuracy • Radiation beam width • High-contrast resolution • Low-contrast resolution • Geometric or distance accuracy • CT number accuracy and uniformity • Artifact evaluation | D: Documentation is required; 3: Direct impact requirements apply | CT
A 20 | At least annually, a diagnostic medical physicist or magnetic resonance imaging (MRI) scientist conducts a performance evaluation of all MRI imaging equipment. The evaluation results, along with recommendations for correcting any problems identified, are documented. The evaluation includes the use of phantoms to assess the following imaging metrics: 5 • Image uniformity for all radiofrequency (RF) coils used clinically • Signal-to-noise ratio (SNR) for all coils used clinically • Slice thickness accuracy • Slice position accuracy • Alignment light accuracy • High-contrast resolution • Low-contrast resolution (or contrast-to-noise ratio) • Geometric or distance accuracy • Magnetic field homogeneity • Artifact evaluation | D: Documentation is required; 3: Direct impact requirements apply | MRI
A 21 | At least annually, a diagnostic medical physicist or nuclear medicine physicist conducts a performance evaluation of all nuclear medicine imaging equipment. The evaluation results, along with recommendations for correcting any problems identified, are documented. The evaluations are conducted for all of the image types produced clinically by each NM scanner (for example, planar and/or tomographic) and include the use of phantoms to assess the following imaging metrics: 6 7 • Image uniformity/system uniformity • High-contrast resolution/system spatial resolution • Sensitivity • Energy resolution • Count-rate performance • Artifact evaluation | D: Documentation is required; 3: Direct impact requirements apply | NM
A 22 | At least annually, a diagnostic medical physicist conducts a performance evaluation of all positron emission tomography (PET) imaging equipment. The evaluation results, along with recommendations for correcting any problems identified, are documented. The evaluations are conducted for all of the image types produced clinically by each PET scanner (for example, planar and/or tomographic) and include the use of phantoms to assess the following imaging metrics: 4 8 • Image uniformity/system uniformity • High-contrast resolution/system spatial resolution • Low-contrast resolution or detectability (not applicable for planar acquisitions) • Artifact evaluation | D: Documentation is required; 3: Direct impact requirements apply | PET
A 23 | For computed tomography (CT), positron emission tomography (PET), nuclear medicine (NM), or magnetic resonance imaging (MRI) services: The annual performance evaluation conducted by the diagnostic medical physicist or MRI scientist (for MRI only) includes testing of image acquisition display monitors for maximum and minimum luminance, luminance uniformity, resolution, and spatial accuracy. 2 5 |  | CT, MRI, PET, NM
EC.02.06.05 | The [critical access] hospital manages its environment during demolition, renovation, or new construction to reduce risk to those in the organization.
A 4 | For computed tomography (CT), positron emission tomography (PET), or nuclear medicine (NM) services: Prior to installation of new imaging equipment, replacement of existing imaging equipment, or modification to rooms where ionizing radiation will be emitted or radioactive materials will be stored (such as scan rooms or hot labs), a medical physicist or health physicist conducts a structural shielding design 9 assessment to specify required radiation shielding. Note: This element of performance does not apply to dental cone beam CT radiographic imaging studies performed for diagnosis of conditions affecting the maxillofacial region or to obtain guidance for the treatment of such conditions. |  | CT, PET, NM
A 6 | For computed tomography (CT), positron emission tomography (PET), or nuclear medicine (NM) services: After installation of imaging equipment or construction in rooms where ionizing radiation will be emitted or radioactive materials will be stored, a medical physicist or health physicist conducts a radiation protection survey to verify the adequacy of installed shielding.9 This survey is conducted prior to clinical use of the room. Note: This element of performance does not apply to dental cone beam CT radiographic imaging studies performed for diagnosis of conditions affecting the maxillofacial region or to obtain guidance for the treatment of such conditions. |  | CT, PET, NM
Human Resources (HR)
HR.01.02.05 | The [critical access] hospital verifies staff qualifications.
C 20 | The [critical access] hospital verifies and documents that diagnostic medical physicists who support computed tomography (CT) services have board certification in diagnostic radiologic physics or radiologic physics by the American Board of Radiology, or in Diagnostic Imaging Physics by the American Board of Medical Physics, or in Diagnostic Radiological Physics by the Canadian College of Physicists in Medicine, or meet all of the following requirements: 2 • A graduate degree in physics, medical physics, biophysics, radiologic physics, medical health physics, or a closely related science or engineering discipline from an accredited college or university • College coursework in the biological sciences with at least one course in biology or radiation biology and one course in anatomy, physiology, or a similar topic related to the practice of medical physics • Documented experience in a clinical CT environment conducting at least 10 CT performance evaluations under the direct supervision of a board-certified medical physicist. | D: Documentation is required; M: Measure of Success is needed | CT
HR.01.05.03 | Staff participate in ongoing education and training.
C 14 | The [critical access] hospital verifies and documents that technologists who perform diagnostic computed tomography (CT) examinations participate in ongoing education that includes annual training on the following: 2 10 11 • Radiation dose optimization techniques and tools for pediatric and adult patients addressed in the Image Gently® and Image Wisely® campaigns • Safe procedures for operation of the types of CT equipment they will use | D: Documentation is required; M: Measure of Success is needed | CT
C 25 | The [critical access] hospital verifies and documents that technologists who perform magnetic resonance imaging (MRI) examinations participate in ongoing education that includes annual training on safe MRI practices in the MRI environment, including the following: • Patient screening criteria that address ferromagnetic items, electrically conductive items, medical implants and devices, and risk for Nephrogenic Systemic Fibrosis (NSF) • Proper patient and equipment positioning activities to avoid thermal injuries • Equipment and supplies that have been determined to be acceptable for use in the MRI environment (MR safe or MR conditional) 12 • MRI safety response procedures for patients who require urgent or emergent medical care • MRI system emergency shutdown procedures, such as MRI system quench and cryogen safety procedures • Patient hearing protection • Management of patients with claustrophobia, anxiety, or emotional distress | D: Documentation is required; M: Measure of Success is needed | MRI
Medication Management (MM)
MM.06.01.01 | The [critical access] hospital safely administers medications.
C 13 | Before administering a radioactive pharmaceutical for diagnostic purposes, staff verify that the dose to be administered is within 20% of the prescribed dose, or, if the dose is prescribed as a range, staff verify that the dose to be administered is within the prescribed range. | 3: Direct impact requirements apply; M: Measure of Success is needed | PET, NM
Provision of Care, Treatment, and Services (PC)
PC.01.02.15 | The [critical access] hospital provides for diagnostic testing.
C 5 | The [critical access] hospital documents in the patient’s record the radiation dose index (CTDIvol, DLP, or size-specific dose estimate [SSDE]) on every study produced during a diagnostic computed tomography (CT) examination. The radiation dose index must be exam specific, summarized by series or anatomic area, and documented in a retrievable format. 2 11 13 15 | D: Documentation is required; 3: Direct impact requirements apply; M: Measure of Success is needed | CT
A 10 | For [critical access] hospitals that provide diagnostic computed tomography (CT), magnetic resonance imaging (MRI), positron emission tomography (PET), or nuclear medicine (NM) services: Prior to conducting a diagnostic imaging study, the [critical access] hospital verifies the following: 2 • Correct patient • Correct imaging site • Correct patient positioning • For CT only: Correct imaging protocol • For CT only: Correct scanner parameters | 3: Direct impact requirements apply | CT, MRI, PET, NM
A 12 | For [critical access] hospitals that provide diagnostic computed tomography (CT), magnetic resonance imaging (MRI), positron emission tomography (PET), or nuclear medicine (NM) services: The [critical access] hospital considers the patient’s age and recent imaging exams when deciding on the most appropriate type of imaging exam. 2 16 | 3: Direct impact requirements apply | CT, MRI, PET, NM
PC.01.03.01 | The [critical access] hospital plans the patient’s care.
A 25 | The [critical access] hospital establishes or adopts diagnostic computed tomography (CT) imaging protocols based on current standards of practice, which address key criteria including clinical indication, contrast administration, age (to indicate whether the patient is pediatric or an adult), patient size and body habitus, and the expected radiation dose index range. 2 | 3: Direct impact requirements apply | CT
A 26 | Diagnostic computed tomography (CT) imaging protocols are reviewed and kept current with input from an interpreting radiologist, medical physicist, and lead imaging technologist to make certain that they adhere to current standards of practice and account for changes in CT imaging equipment. These reviews are conducted at time frames identified by the [critical access] hospital. 2 | 3: Direct impact requirements apply | CT
Performance Improvement (PI)
PI.01.01.01 | The [critical access] hospital collects data to monitor its performance.
A 46 | The [critical access] hospital collects data on patient thermal injuries that occur during magnetic resonance imaging exams. |  | MRI
A 47 | The [critical access] hospital collects data on the following: • Incidents where ferromagnetic objects unintentionally entered the magnetic resonance imaging (MRI) scanner room • Injuries resulting from the presence of ferromagnetic objects in the MRI scanner room |  | MRI
PI.02.01.01 | The [critical access] hospital compiles and analyzes data.
A 6 | The [critical access] hospital reviews and analyzes incidents where the radiation dose index (CTDIvol, DLP, or size-specific dose estimate [SSDE]) from diagnostic CT examinations exceeded expected dose index ranges identified in imaging protocols. These incidents are then compared to external benchmarks. 2 15 |  | CT
Notes:
1 For the definition of ALARA, please refer to U.S. Nuclear Regulatory Commission federal regulation 10 CFR 20.1003. 2 This element of performance does not apply to dental cone beam CT radiographic imaging studies performed for diagnosis of conditions affecting the maxillofacial region or to obtain guidance for the treatment of such conditions. 3 This element of performance is only applicable for systems capable of calculating and displaying radiation doses. 4 Medical physicists are accountable for these activities. They may be assisted with the testing and evaluation of equipment performance by individuals who have the required training and skills, as determined by the physicist. (See also HR.01.02.01, EP 1; HR.01.02.05, EP 20; HR.01.02.07, EPs 1 and 2; HR.01.06.01, EP 1; and LD.03.06.01, EP 4) 5 Medical physicists or MRI scientists are accountable for these activities. They may be assisted with the testing and evaluation of equipment performance by individuals who have the required training and skills, as determined by the medical physicist or MRI scientist. (See also HR.01.02.01, EP 1; HR.01.02.05, EP 20; HR.01.02.07, EPs 1 and 2; HR.01.06.01, EP 1; and LD.03.06.01, EP 4) 6 The following test is recommended, but not required: Low-contrast resolution or detectability for nonplanar acquisitions. 7 The medical physicist or nuclear medicine physicist is accountable for these activities. He or she may be assisted with the testing and evaluation of equipment performance by individuals who have the required training and skills, as determined by the medical physicist or nuclear medicine physicist. (See also HR.01.02.01, EP 1; HR.01.02.05, EP 20; HR.01.02.07, EPs 1 and 2; HR.01.06.01, EP 1; and LD.03.06.01, EP 4) 8 The following tests are recommended, but not required, for PET scanner testing: sensitivity, energy resolution, and count-rate performance. 9 For additional guidance on shielding designs and radiation protection surveys, see National Council on Radiation Protection and Measurements Report No. 147 (NCRP-147). 10 Information on the Image Gently and Image Wisely initiatives can be found online at http://www.imagegently.org and http://www.imagewisely.org, respectively. 11 This element of performance does not apply to CT systems used for therapeutic radiation treatment planning or delivery, or for calculating attenuation coefficients for nuclear medicine studies. 12 Terminology for defining the safety of items in the magnetic resonance environment is provided in ASTM F2503 Standard Practice for Marking Medical Devices and Other Items for Safety in the Magnetic Resonance Environment (http://www.astm.org). 13 This element of performance is only applicable for systems capable of calculating and displaying radiation dose indices. 15 While the CTDIvol, DLP, and SSDE are useful indicators for monitoring radiation dose indices from the CT machine, they do not represent the patient’s radiation dose. 16 Knowledge of a patient’s recent imaging exams can help to prevent unnecessary duplication of these examinations.
The Joint Commission Prepublication Requirements for Diagnostic Imaging Services (issued January 9, 2015, effective July 1, 2015)
http://www.jointcommission.org/assets/1/6/HAP-CAH_DiagImag_Prepub_July2015release_20150105.pdf
Note on C9: 1 For the definition of ALARA, please refer to U.S. Nuclear Regulatory Commission federal regulation 10 CFR 20.1003. 2 This element of performance does not apply to dental cone beam CT radiographic imaging studies performed for diagnosis of conditions affecting the maxillofacial region or to obtain guidance for the treatment of such conditions.
Note on C14: 2 This element of performance does not apply to dental cone beam CT radiographic imaging studies performed for diagnosis of conditions affecting the axillofacial region or to obtain guidance for the treatment of such conditions. 3 This element of performance is only applicable for systems capable of calculating and displaying radiation doses. 4 Medical physicists are accountable for these activities. They may be assisted with the testing and evaluation of equipment performance by individuals who have the required training and skills, as determined by the physicist. (See also HR.01.02.01, EP 1; HR.01.02.05, EP 20; HR.01.02.07, EPs 1 and 2; HR.01.06.01, EP 1; and LD.03.06.01, EP 4)
Note on C15: 2 This element of performance does not apply to dental cone beam CT radiographic imaging studies performed for diagnosis of conditions affecting the maxillofacial region or to obtain guidance for the treatment of such conditions. 4 Medical physicists are accountable for these activities. They may be assisted with the testing and evaluation of equipment performance by individuals who have the required training and skills, as determined by the physicist. (See also HR.01.02.01, EP 1; HR.01.02.05, EP 20; HR.01.02.07, EPs 1 and 2; HR.01.06.01, EP 1; and LD.03.06.01, EP 4)
Note on C16: 5 Medical physicists or MRI scientists are accountable for these activities. They may be assisted with the testing and evaluation of equipment performance by individuals who have the required training and skills, as determined by the medical physicist or MRI scientist. (See also HR.01.02.01, EP 1; HR.01.02.05, EP 20; HR.01.02.07, EPs 1 and 2; HR.01.06.01, EP 1; and LD.03.06.01, EP 4)
Note on C17: 6 The following test is recommended, but not required: Low-contrast resolution or detectability for nonplanar acquisitions. 7 The medical physicist or nuclear medicine physicist is accountable for these activities. He or she may be assisted with the testing and evaluation of equipment performance by individuals who have the required training and skills, as determined by the medical physicist or nuclear medicine physicist. (See also HR.01.02.01, EP 1; HR.01.02.05, EP 20; HR.01.02.07, EPs 1 and 2; HR.01.06.01, EP 1; and LD.03.06.01, EP 4)
Note on C18: 4 Medical physicists are accountable for these activities. They may be assisted with the testing and evaluation of equipment performance by individuals who have the required training and skills, as determined by the medical physicist. (See also HR.01.02.01, EP 1; HR.01.02.05, EP 20; HR.01.02.07, EPs 1 and 2; HR.01.06.01, EP 1; and LD.03.06.01, EP 4) 8 The following tests are recommended, but not required, for PET scanner testing: sensitivity, energy resolution, and count-rate performance.
Note on C19: 2 This element of performance does not apply to dental cone beam CT radiographic imaging studies performed for diagnosis of conditions affecting the maxillofacial region or to obtain guidance for the treatment of such conditions. 5 Medical physicists or MRI scientists are accountable for these activities. They may be assisted with the testing and evaluation of equipment performance by individuals who have the required training and skills, as determined by the physicist or MRI scientist. (See also HR.01.02.01, EP 1; HR.01.02.05, EP 20; HR.01.02.07, EPs 1 and 2; HR.01.06.01, EP 1; and LD.03.06.01, EP 4)
Note on C21: 9 For additional guidance on shielding designs and radiation protection surveys, see National Council on Radiation Protection and Measurements Report No. 147 (NCRP-147).
Note on C22: 9 For additional guidance on shielding designs and radiation protection surveys, see National Council on Radiation Protection and Measurements Report No. 147 (NCRP-147).
Note on C25: 2 This element of performance does not apply to dental cone beam CT radiographic imaging studies performed for diagnosis of conditions affecting the maxillofacial region or to obtain guidance for the treatment of such conditions.
Note on C27: 2 This element of performance does not apply to dental cone beam CT radiographic imaging studies performed for diagnosis of conditions affecting the maxillofacial region or to obtain guidance for the treatment of such conditions. 10 Information on the Image Gently and Image Wisely initiatives can be found online at http://www.imagegently.org and http://www.imagewisely.org, respectively. 11 This element of performance does not apply to CT systems used for therapeutic radiation treatment planning or delivery, or for calculating attenuation coefficients for nuclear medicine studies.
Note on C28: 12 Terminology for defining the safety of items in the magnetic resonance environment is provided in ASTM F2503 Standard Practice for Marking Medical Devices and Other Items for Safety in the Magnetic Resonance Environment (http://www.astm.org).
Note on C34: 2 This element of performance does not apply to dental cone beam CT radiographic imaging studies performed for diagnosis of conditions affecting the maxillofacial region or to obtain guidance for the treatment of such conditions. 11 This element of performance does not apply to systems used for therapeutic radiation treatment planning or delivery, or for calculating attenuation coefficients for nuclear medicine studies. 13 This element of performance is only applicable for systems capable of calculating and displaying radiation dose indices. 15 While the CTDIvol, DLP, and SSDE are useful indicators for monitoring radiation dose indices from the CT machine, they do not represent the patient’s radiation dose.
Note on C35: 2 This element of performance does not apply to dental cone beam CT radiographic imaging studies performed for diagnosis of conditions affecting the maxillofacial region or to obtain guidance for the treatment of such conditions.
Note on C36: 2 This element of performance does not apply to dental cone beam CT radiographic imaging studies performed for diagnosis of conditions affecting the maxillofacial region or to obtain guidance for the treatment of such conditions. 16 Knowledge of a patient’s recent imaging exams can help to prevent unnecessary duplication of these examinations.
Note on C38: 2 This element of performance does not apply to dental cone beam CT radiographic imaging studies performed for diagnosis of conditions affecting the maxillofacial region or to obtain guidance for the treatment of such conditions.
Note on C39: 2 This element of performance does not apply to dental cone beam CT radiographic imaging studies performed for diagnosis of conditions affecting the maxillofacial region or to obtain guidance for the treatment of such conditions.
Note on C45: 2 This element of performance does not apply to dental cone beam CT radiographic imaging studies performed for diagnosis of conditions affecting the maxillofacial region or to obtain guidance for the treatment of such conditions. 15 While the CTDIvol, DLP, and SSDE are useful indicators for monitoring radiation dose indices from the CT machine, they do not represent the patient’s radiation dose.

## Summary
The Joint Commission Revised Standards Gap Analysis Executive Summary
Date: _
Summary will be generated after gap analysis. Start typing in the 'Gap Analysis' tab!
| Status |  |  |  | Summary
| Not in place | Partially met | Fully met | N/A | Not in place | Partially met | Fully met | N/A
Environment of Care (EC) |  |  |  |  | 0 | 0 | 0 | 0 | #DIV/0! | #DIV/0! | #DIV/0!
EC.02.01.01.A14
EC.02.01.01.A16
EC.02.02.01.A17
EC.02.04.01.A10
EC.02.04.03.C15
EC.02.04.03.A17
EC.02.04.03.A19
EC.02.04.03.A20
EC.02.04.03.A21
EC.02.04.03.A22
EC.02.04.03.A23
EC.02.06.05.A4
EC.02.06.05.A6
Human Resources (HR)
HR.01.02.05.C20
HR.01.05.03.C14
HR.01.05.03.C25
Medication Management (MM)
MM.06.01.01.C13
Provision of Care, Treatment, and Services (PC)
PC.01.02.15.C5
PC.01.02.15.A10
PC.01.02.15.A12
PC.01.03.01.A25
PC.01.03.01.A26
Performance Improvement (PI)
PI.01.01.01.A46
PI.01.01.01.A47
PI.02.01.01.A6